상 세 규 격 서
⃞ 장비명 : 무균대
구 분	상 세 명 세	수 량	비 고
1. Feature : 2. Specification : ① Main System ② Option ③ 설치조건 ④ 보증기간 납품기한 등	1. The intent herein is to provide a concise statement of requirements for a quality Class II Laminar Airflow Biological Safety Cabinet 2. HEPEX Zero Leak Airflow System The HEPEX Zero Leak Airflow System eliminates the possibility of cabinet and gasket leaks 3. Greater Volume of Air Creates a stronger Air Barrier Air exchanges within the ClassⅡ Type A2 Biological Safety Cabinet occur 25 times per minute 4. Larger, Long Lasting HEPA Filters HEPA filters that are typically 25% larger than those in competitive products. 5. Motor/Blower System Extends HEPA Filter Life As the HEPA filter loads the motor/blower system automatically compensates to assure performance and extended filter life. 6. The instrument meets the performance requirements of NSF/ANSI 49:2007. Your confidence is well placed in a Biological Safety Cabinet that meets NSF Standard. 7. The cabinet shall have all positive pressure plenums surrounded by a vacuum relative to the room. 8. The cabinet shall have separate internal electrical circuits for blower, fluorescent light, each outlet, accessory connection and optional UV light. Each circuit shall be protected with afuse located in the Control Center on the electronic module. 9. The cabinet shall have two internal electrical circuits; one for blower/lights and one for the duplex outlets. Each circuit shall be protected with a circuit breaker located in the Control Center. 10. Total cabinet airflow shall be controlled via a solid-state motor voltage regulator with adjustment available on the Control Center. 11. Balancing of cabinet work zone downflow (recycling flow) to exhaust flow shall be accomplished with an internal exhaust flow damper, externally adjustable with screwdriver and sealed with liquid tight fastener. 12. The cabinet shall be easily transportable through a standard 36 inch wide door without disassembly. 13. Fluorescent lighting shall be externally mounted and provide 90 (968) to 120 (1291) footcandles (LUX) on work surface. The ballast is to be electronic containing thermal protection with automatic reset. 14. Cabinet shall come standard with two outlets with drip proof covers on back wall ; one gas valve/service coupling on right side wall and one service coupling on right side wall. 15. Cabinet work zone shall be all 16 GA. stainless steel and reinforced with stainless steel U channels to minimize vibration. 16. A 3/8 inch ball valve with welded fitting shall be provided in the drain trough beneath the work tray 17. Cabinet shall have a permanent positive pressure plenum with quick release supply filter removal. 18. Cabinet shall be capable of front filter removal without disassembly of the control panel and sliding window tracks/hardware. 19. Harnessing the efficiency of DC ECM Technology : Filter Loading Capacity 10 Years 1. Type : Class II A2 Type ( Recirculated Air 70%, Exhausted Air 30%) 2. Metal Framed HEPA Filter : Minimum filter efficiency is 99.99% for particulates 0.3 micrometers. Both exhaust and supply HEPA filters can be inserted and removed from the front of the cabinet. 3. Interior Dimensions (mm) : 1,788 W x 597 D x 724 H or more 4. Overall Dimensions (mm): 1,972 W x 835 D x 1,600 H or less 5. Adjustable Support Stand height : 673~ 825mm of adjustable 6. Cabinet's material & structure : 16gauge, pressure tight, all welded, Stainless steel 7. Control System : Microprocessor Control 8. Display : Innovativeelectronic controller includes and diagnostic functions. 9. Window : Tempered Glass Window 10. Access opening : 254mm 11. Air Velocity from the supply filter shall average 55 to 65 FPM with no single point outside the 20% of average range measured in a horizontal plane defined by 4 inches above the bottom edge of window. 12. Work access opening shall be 8 inches high. Average inflow velocity shall nominally be 105 LFPM. 13. Sound level shall be no more than 63 dbA measured 15 inches above the work tray and 12 inches in front of viewing window. 14. Motor/blower shall be positioned so as to create an even filter loading, thereby prolonging the life of the supply HEPA filter, and shall deliver over 80% the initial HEPA filter static pressure with no more than a 10% decrease of CFM. 1. Adjustable Support Stand : 1ea 2. UV Light(Include Spare Part) : 1 ea 3. Armrest: 1ea 4. Electric Outlet : 2ea 5. Gas Valve : 1ea 6. Operating Manual : 1copy 7. Service Manual : 1copy 1. 본연구원 실험실(7층) 으로 안전하게 배송 설치 해야하며 추후 본건물로 이전설치도 보장되어야함. 1. The equipment should be installed by supplier. 2. Three year warranty after the performance test.	1 set
2013 년 9월 16 일 부서장(연구책임자) : 허향숙